# Patient-Level Normalization for ECG-Based
Haemodynamic-Instability Classification

Frozen Results, Methodological Audit, and Implications for Patient Adaptation

Internal Collaborator Results Report

20-patient leave-one-subject-out validation

Analysis frozen | 31 July 2026

This report reconstructs and interprets the supplied frozen evidence. No model was retrained, tuned, or replaced, and no alternative result was introduced.

Prepared for internal scientific collaboration and manuscript development

# Contents

1. Purpose and Intended Audience

2. Executive Summary

3. Clinical and Scientific Motivation

4. Study Objectives and Research Questions

5. Dataset and Analysis Unit

6. Experimental Design

7. Patient-Level Normalization Strategies

8. Leakage-Safe LOSO Validation Pipeline

9. Feature Domains and Fold-Specific Feature Selection

10. Classifiers and Model Fitting

11. Evaluation Populations and Performance Measures

12. Reproducibility and Audit Controls

13. Frozen Results

14. Patient-Level Heterogeneity

15. Calibration and IQR Fallback Findings

16. Main Scientific Interpretation

17. What the Study Does and Does Not Establish

18. Limitations

19. Implications for a Practical Monitoring System

20. Proposed Manuscript Story

21. Candidate Manuscript Contributions

22. Recommended Tables and Figures

23. Conclusions

24. Technical Appendix

In Microsoft Word, right-click this contents page and select Update Field if an automatic page-numbered table of contents is desired.

# 1. Purpose and Intended Audience

This report provides collaborators with a complete, auditable account of the frozen ECG-based haemodynamic-instability classification study. It reconstructs the analytical design, explains the normalization and validation pipeline, consolidates pooled and patient-level findings, and separates verified results from interpretation, limitations, and future work.

The report is written for biomedical researchers, clinicians, data scientists, and manuscript co-authors. Python implementation details are included only where they materially affect scientific validity, leakage control, or reproducibility.

The repository begins with an already constructed table of ECG-derived features. The upstream waveform retrieval, event labelling, segmentation, signal preprocessing, and feature-extraction pipeline are outside the minimal frozen repository and cannot be independently audited from this package.

# 2. Executive Summary

The study evaluated whether patient-level normalization improves ECG-based classification of haemodynamic instability in 20 patients. The frozen dataset contained 1,260 ECG-derived observations: 593 reference observations and 667 HI observations. Twenty-fold leave-one-subject-out validation compared four normalization methods, two feature domains, and three classifiers, producing 24 completed configurations.

The primary comparison used the same 1,060 non-calibration rows for every method. This common-row design holds patient composition, row identity, sample size, and prevalence constant, making it the most appropriate method-to-method evaluation.

Without normalization, tree models often showed much stronger within-patient discrimination than pooled discrimination. Reduced-feature Random Forest had pooled AUC 0.639 but median patient AUC 0.849, suggesting cross-patient score misalignment.

Retrospective jackknife normalization produced the strongest tree-model result: reduced-feature Random Forest AUC 0.866, average precision 0.889, Brier score 0.179, and median patient AUC 0.966. Sixteen patients improved relative to no normalization.

Fixed 10-row labelled calibration recovered part of the retrospective benefit. The best calibration-median AUC was 0.728, while the best calibration median-IQR AUC was 0.769.

For tree models, the retrospective-to-calibration AUC gap ranged from 0.064 to 0.169, depending on feature domain and calibration method.

Median-IQR scaling improved pooled tree-model AUC over median centering by 0.032-0.052, but patient-level effects were mixed and fallback was required in 18% of eligible applications.

Normalization effects were classifier-dependent: Extra Trees and Random Forest benefited substantially, whereas the fixed Logistic Regression implementation did not.

Central conclusion: Patient adaptation can materially improve ECG-based HI classification, but its apparent value depends on the quantity and timing of patient-specific information, the classifier, the feature representation, and whether evaluation is pooled or patient-level.

# 3. Clinical and Scientific Motivation

Continuous invasive blood-pressure monitoring is not available for every ICU patient or throughout every admission. ECG is among the most routinely acquired continuous physiological waveforms. Although ECG does not directly measure blood pressure or cardiac output, its morphology and dynamics are influenced by rhythm, conduction, autonomic regulation, myocardial function, medications, devices, and systemic stress. This motivates testing whether ECG-derived features contain information associated with haemodynamic instability.

Patient-independent ECG modelling is difficult because baseline feature distributions differ between patients. A classifier may distinguish reference and HI observations within one patient while assigning poorly aligned absolute scores across different patients. Patient normalization attempts to reduce these stable patient-specific offsets by representing each observation relative to a patient-specific reference centre or scale.

The scientific meaning of normalization depends on the information used. Retrospective use of a patient’s labelled reference record demonstrates potential under favourable information availability. Fixed limited calibration asks how much of that benefit remains when personalization is based on only 10 labelled reference observations.

# 4. Study Objectives and Research Questions

Primary objective: determine whether patient-level normalization improves subject-independent classification of ECG observations associated with haemodynamic instability under leave-one-subject-out validation.

RQ1. How well do engineered ECG features discriminate reference and HI observations in unseen patients without personalization?

RQ2. Does the difference between pooled and patient-level AUC indicate inter-patient score misalignment?

RQ3. How much does retrospective jackknife median normalization improve classification?

RQ4. How much of that benefit is retained using only 10 labelled calibration observations?

RQ5. Does median-IQR scaling improve beyond median centering?

RQ6. Do effects differ between Logistic Regression, Extra Trees, and Random Forest?

RQ7. Do effects differ between the all-feature and reduced feature domains?

RQ8. Are aggregate gains broadly distributed across patients?

RQ9. How frequently does IQR scaling require training-derived fallback?

RQ10. Do conclusions remain stable between common-row and native-row evaluation?

## Table 1. Dataset and Study Design

| Item | Frozen value |
| --- | --- |
| Patients | 20 |
| ECG-derived observations | 1,260 |
| Reference/Class 0 observations | 593 |
| HI/Class 1 observations | 667 |
| Validation | 20-fold LOSO |
| Normalization methods | 4 |
| Feature domains | 2 |
| Classifiers | 3 |
| Configurations | 24 |
| Calibration observations | 10 per patient |
| Random seed | 42 |

# 5. Dataset and Analysis Unit

The canonical modelling file was data/FullFeatureDataset_Clean.csv. It contained 35 candidate ECG features and three metadata columns: patient_id, segment_id, and Class. Every patient contributed both classes, allowing patient-level ROC AUC calculation in all 20 held-out folds.

The analysis unit was one ECG-derived observation represented by one feature-table row. The 1,260 rows are repeated observations nested within 20 patients; they are not 1,260 independent subjects. Holding out entire patients is therefore essential.

Audit checks found no exact duplicate rows, missing feature values, infinite feature values, or non-numeric candidate features. The patient_id + segment_id combination was not unique, so segment identifiers and deterministic row ordering cannot be treated as verified clinical chronology.

Canonical dataset SHA-256: ef5a8c73034afa0ff92da1a5c31d133c1d847fe1894ba8c559bfd2cdab440796.

# 6. Experimental Design

The factorial benchmark crossed four normalization methods with two feature domains and three classifiers. In each of 20 outer folds, all observations from one patient were held out, while all train-derived preprocessing and feature selection were fitted on the remaining 19 patients. No inner cross-validation or hyperparameter search was performed; the design is therefore LOSO validation with fold-internal preprocessing and feature selection, not nested cross-validation.

Figure 1. Leakage-safe LOSO pipeline. The held-out patient is excluded from train-derived preprocessing, feature selection, scaling, and model fitting.

# 7. Patient-Level Normalization Strategies

## 7.1 No normalization

Original feature values were retained. No information from the held-out patient was used to personalize the representation.

## 7.2 Retrospective jackknife median

Each observation was centred using labelled reference observations from the same patient. For a reference row, the current row was excluded from its centring median. This is a retrospective, optimistic comparator and is not a directly deployable initial-calibration method.

## 7.3 Fixed labelled calibration median

Ten labelled Class 0 observations per patient were used to estimate a fixed patient median. The 200 calibration observations were excluded from fitting and evaluation, leaving 1,060 common rows. Selection was deterministic but not verified as chronological.

## 7.4 Fixed labelled calibration median-IQR

The same 10 observations were used to estimate median and IQR. If the patient IQR was zero, non-finite, or below tolerance, the implementation used a training-derived fallback: median non-zero outer-training patient IQR, then pooled outer-training calibration IQR, then scale 1.0.

Figure 2. Conceptual comparison of the four normalization strategies and the amount of held-out-patient information used.

# 8. Leakage-Safe LOSO Validation Pipeline

The frozen order was preserved exactly: patient normalization → calibration-row exclusion → infinity handling → training-only median imputation → variance filtering → Random Forest feature ranking → top-20 selection → Spearman correlation pruning at 0.90 → training-only StandardScaler → classifier fitting → held-out-patient prediction.

The implementation includes checks for train-test row overlap, patient overlap, calibration-row exclusion, jackknife self-exclusion, held-out-patient contribution to IQR fallback, and consistency of common-row identities. Fold-level hashes of transformed train and test matrices were retained.

# 9. Feature Domains and Fold-Specific Feature Selection

The complete domain contained 35 features spanning time/amplitude/distribution, entropy, nonlinear complexity, fractal, and frequency-domain summaries. The reduced domain contained 10 features: mean absolute value, median, minimum, maximum, variance, standard deviation, skewness, kurtosis, positive-to-negative sample ratio, and number of zero crossings.

Feature selection was repeated independently inside every LOSO training fold. Features were ranked by a training-only Random Forest, restricted to the top 20, and pruned for Spearman correlation above 0.90. Consequently, the benchmark evaluates a fold-specific selection pipeline rather than one globally fixed biomarker set.

# 10. Classifiers and Model Fitting

Logistic Regression represented a linear classifier. Extra Trees and Random Forest represented nonlinear tree ensembles. Both ensembles used 500 trees, balanced class weighting, and seed 42. Model specifications were fixed; no hyperparameter tuning, threshold optimization, or inner model-selection loop was performed.

# 11. Evaluation Populations and Performance Measures

The primary common-row population contained the same 1,060 observations for all configurations: 393 reference and 667 HI observations. Native-row evaluation retained 1,260 rows for no normalization and jackknife, but 1,060 rows for the calibration methods. Native-row comparisons are secondary because row populations and positive prevalence differ.

Primary discrimination was assessed with pooled ROC AUC, complemented by average precision, Brier score, balanced accuracy, and patient-level ROC AUC. Patient direction categories identified correctly directed, near-random, and inverted discrimination. Pooled and patient-level metrics answer different questions: pooled AUC includes cross-patient score alignment, whereas patient AUC measures within-patient ranking.

# 12. Reproducibility and Audit Controls

The frozen package retained a manifest, row-level out-of-fold predictions, common and native pooled metrics, patient-level metrics, directionality summaries, calibration rows, fallback summaries, fold summaries, selected features, and transformed-matrix hashes. The OOF table contained 30,240 records, exactly 1,260 rows × 24 configurations. The 2,400 missing probabilities corresponded exactly to 200 excluded calibration rows × 12 calibration configurations.

The manifest recorded Python 3.13.9, NumPy 2.3.5, pandas 2.3.3, and scikit-learn 1.7.2. It also recorded a non-clean Git working tree. The archive did not include .git history, and a complete exact-environment numerical rerun was not performed during this report assembly.

# Table 2. Complete Experimental Matrix

| Normalization | Feature domain | Classifier | Folds | Common rows | Native rows |
| --- | --- | --- | --- | --- | --- |
| No normalization | All features | Logistic Regression | 20 | 1060 | 1260 |
| No normalization | All features | Extra Trees | 20 | 1060 | 1260 |
| No normalization | All features | Random Forest | 20 | 1060 | 1260 |
| No normalization | Reduced features | Logistic Regression | 20 | 1060 | 1260 |
| No normalization | Reduced features | Extra Trees | 20 | 1060 | 1260 |
| No normalization | Reduced features | Random Forest | 20 | 1060 | 1260 |
| Retrospective jackknife | All features | Logistic Regression | 20 | 1060 | 1260 |
| Retrospective jackknife | All features | Extra Trees | 20 | 1060 | 1260 |
| Retrospective jackknife | All features | Random Forest | 20 | 1060 | 1260 |
| Retrospective jackknife | Reduced features | Logistic Regression | 20 | 1060 | 1260 |
| Retrospective jackknife | Reduced features | Extra Trees | 20 | 1060 | 1260 |
| Retrospective jackknife | Reduced features | Random Forest | 20 | 1060 | 1260 |
| Calibration median | All features | Logistic Regression | 20 | 1060 | 1060 |
| Calibration median | All features | Extra Trees | 20 | 1060 | 1060 |
| Calibration median | All features | Random Forest | 20 | 1060 | 1060 |
| Calibration median | Reduced features | Logistic Regression | 20 | 1060 | 1060 |
| Calibration median | Reduced features | Extra Trees | 20 | 1060 | 1060 |
| Calibration median | Reduced features | Random Forest | 20 | 1060 | 1060 |
| Calibration median-IQR | All features | Logistic Regression | 20 | 1060 | 1060 |
| Calibration median-IQR | All features | Extra Trees | 20 | 1060 | 1060 |
| Calibration median-IQR | All features | Random Forest | 20 | 1060 | 1060 |
| Calibration median-IQR | Reduced features | Logistic Regression | 20 | 1060 | 1060 |
| Calibration median-IQR | Reduced features | Extra Trees | 20 | 1060 | 1060 |
| Calibration median-IQR | Reduced features | Random Forest | 20 | 1060 | 1060 |

# 13. Frozen Results

## 13.1 Performance without normalization

Without normalization, Logistic Regression remained near chance (AUC 0.474 with all features; 0.515 with reduced features). Tree models performed better with the reduced domain: Extra Trees AUC 0.637 and Random Forest AUC 0.639. The most important finding was the pooled-versus-patient discrepancy: reduced-feature Random Forest had median patient AUC 0.849 despite pooled AUC 0.639, with six inverted patients.

## 13.2 Retrospective jackknife normalization

Jackknife centring substantially improved both tree ensembles. The best frozen configuration was reduced-feature Random Forest: AUC 0.866, average precision 0.889, Brier score 0.179, balanced accuracy 0.767, and median patient AUC 0.966. Sixteen patients improved, three worsened, and one was unchanged relative to no normalization. Logistic Regression did not improve.

## 13.3 Fixed calibration median

Fixed median calibration recovered part of the tree-model benefit. The strongest result was all-feature Random Forest with AUC 0.728. Patient responses were more heterogeneous: all-feature Random Forest improved 14 patients, worsened five, and left one unchanged; reduced-feature Random Forest split evenly, with 10 improving and 10 worsening.

## 13.4 Fixed calibration median-IQR

Median-IQR scaling improved pooled AUC over median centring for all four tree comparisons by 0.032-0.052. The strongest fixed-calibration AUC was 0.769 with all-feature Extra Trees; all-feature Random Forest achieved AUC 0.760, Brier score 0.200, and balanced accuracy 0.709. Patient-level effects remained mixed.

## 13.5 Classifier and feature-domain interactions

Normalization benefits were concentrated in Extra Trees and Random Forest. The reduced feature domain was stronger without normalization, whereas the complete domain was generally stronger under fixed calibration. No single feature domain or classifier dominated every evaluation objective.

Figure 3. Primary common-row ROC AUC comparison across all 24 configurations. The dashed line marks chance-level ranking.

# Table 3. Primary Common-Row Pooled Results

| Normalization | Domain | Classifier | AUC | AP | Brier | Balanced acc. | Median patient AUC | Correct | Near | Inverted |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Calibration median | All features | Extra Trees | 0.720 | 0.773 | 0.226 | 0.644 | 0.836 | 14 | 1 | 5 |
| Calibration median | All features | Logistic Regression | 0.501 | 0.668 | 0.338 | 0.566 | 0.414 | 7 | 2 | 11 |
| Calibration median | All features | Random Forest | 0.728 | 0.795 | 0.212 | 0.691 | 0.795 | 16 | 0 | 4 |
| Calibration median | Reduced features | Extra Trees | 0.659 | 0.744 | 0.238 | 0.603 | 0.801 | 15 | 3 | 2 |
| Calibration median | Reduced features | Logistic Regression | 0.430 | 0.600 | 0.304 | 0.482 | 0.381 | 8 | 0 | 12 |
| Calibration median | Reduced features | Random Forest | 0.697 | 0.772 | 0.215 | 0.655 | 0.851 | 16 | 2 | 2 |
| Calibration median-IQR | All features | Extra Trees | 0.769 | 0.825 | 0.211 | 0.665 | 0.922 | 18 | 1 | 1 |
| Calibration median-IQR | All features | Logistic Regression | 0.449 | 0.652 | 0.349 | 0.528 | 0.362 | 8 | 0 | 12 |
| Calibration median-IQR | All features | Random Forest | 0.760 | 0.819 | 0.200 | 0.709 | 0.845 | 16 | 2 | 2 |
| Calibration median-IQR | Reduced features | Extra Trees | 0.711 | 0.762 | 0.216 | 0.693 | 0.873 | 15 | 3 | 2 |
| Calibration median-IQR | Reduced features | Logistic Regression | 0.471 | 0.678 | 0.296 | 0.521 | 0.494 | 8 | 2 | 10 |
| Calibration median-IQR | Reduced features | Random Forest | 0.737 | 0.810 | 0.203 | 0.687 | 0.890 | 17 | 2 | 1 |
| Retrospective jackknife | All features | Extra Trees | 0.833 | 0.865 | 0.241 | 0.677 | 0.936 | 17 | 1 | 2 |
| Retrospective jackknife | All features | Logistic Regression | 0.457 | 0.668 | 0.341 | 0.538 | 0.416 | 9 | 0 | 11 |
| Retrospective jackknife | All features | Random Forest | 0.841 | 0.885 | 0.209 | 0.698 | 0.918 | 18 | 0 | 2 |
| Retrospective jackknife | Reduced features | Extra Trees | 0.823 | 0.839 | 0.218 | 0.706 | 0.953 | 19 | 0 | 1 |
| Retrospective jackknife | Reduced features | Logistic Regression | 0.404 | 0.610 | 0.313 | 0.472 | 0.427 | 8 | 1 | 11 |
| Retrospective jackknife | Reduced features | Random Forest | 0.866 | 0.889 | 0.179 | 0.767 | 0.966 | 19 | 1 | 0 |
| No normalization | All features | Extra Trees | 0.566 | 0.644 | 0.261 | 0.509 | 0.682 | 13 | 0 | 7 |
| No normalization | All features | Logistic Regression | 0.474 | 0.619 | 0.302 | 0.475 | 0.495 | 9 | 3 | 8 |
| No normalization | All features | Random Forest | 0.561 | 0.662 | 0.270 | 0.529 | 0.696 | 14 | 3 | 3 |
| No normalization | Reduced features | Extra Trees | 0.637 | 0.746 | 0.236 | 0.600 | 0.831 | 15 | 0 | 5 |
| No normalization | Reduced features | Logistic Regression | 0.515 | 0.675 | 0.267 | 0.489 | 0.471 | 9 | 1 | 10 |
| No normalization | Reduced features | Random Forest | 0.639 | 0.740 | 0.242 | 0.605 | 0.849 | 14 | 0 | 6 |

Note: All 24 methods are evaluated on the same 1,060 rows. This is the primary like-for-like comparison.

# Table 4. Native-Row Pooled Results

| Normalization | Domain | Classifier | Rows | AUC | AP | Brier | Balanced acc. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Calibration median | All features | Extra Trees | 1060 | 0.720 | 0.773 | 0.226 | 0.644 |
| Calibration median | All features | Logistic Regression | 1060 | 0.501 | 0.668 | 0.338 | 0.566 |
| Calibration median | All features | Random Forest | 1060 | 0.728 | 0.795 | 0.212 | 0.691 |
| Calibration median | Reduced features | Extra Trees | 1060 | 0.659 | 0.744 | 0.238 | 0.603 |
| Calibration median | Reduced features | Logistic Regression | 1060 | 0.430 | 0.600 | 0.304 | 0.482 |
| Calibration median | Reduced features | Random Forest | 1060 | 0.697 | 0.772 | 0.215 | 0.655 |
| Calibration median-IQR | All features | Extra Trees | 1060 | 0.769 | 0.825 | 0.211 | 0.665 |
| Calibration median-IQR | All features | Logistic Regression | 1060 | 0.449 | 0.652 | 0.349 | 0.528 |
| Calibration median-IQR | All features | Random Forest | 1060 | 0.760 | 0.819 | 0.200 | 0.709 |
| Calibration median-IQR | Reduced features | Extra Trees | 1060 | 0.711 | 0.762 | 0.216 | 0.693 |
| Calibration median-IQR | Reduced features | Logistic Regression | 1060 | 0.471 | 0.678 | 0.296 | 0.521 |
| Calibration median-IQR | Reduced features | Random Forest | 1060 | 0.737 | 0.810 | 0.203 | 0.687 |
| Retrospective jackknife | All features | Extra Trees | 1260 | 0.812 | 0.781 | 0.219 | 0.673 |
| Retrospective jackknife | All features | Logistic Regression | 1260 | 0.453 | 0.550 | 0.323 | 0.533 |
| Retrospective jackknife | All features | Random Forest | 1260 | 0.827 | 0.823 | 0.194 | 0.698 |
| Retrospective jackknife | Reduced features | Extra Trees | 1260 | 0.795 | 0.748 | 0.206 | 0.687 |
| Retrospective jackknife | Reduced features | Logistic Regression | 1260 | 0.418 | 0.507 | 0.302 | 0.481 |
| Retrospective jackknife | Reduced features | Random Forest | 1260 | 0.842 | 0.821 | 0.175 | 0.748 |
| No normalization | All features | Extra Trees | 1260 | 0.559 | 0.542 | 0.259 | 0.503 |
| No normalization | All features | Logistic Regression | 1260 | 0.482 | 0.522 | 0.301 | 0.487 |
| No normalization | All features | Random Forest | 1260 | 0.556 | 0.563 | 0.265 | 0.517 |
| No normalization | Reduced features | Extra Trees | 1260 | 0.655 | 0.687 | 0.232 | 0.612 |
| No normalization | Reduced features | Logistic Regression | 1260 | 0.525 | 0.590 | 0.261 | 0.497 |
| No normalization | Reduced features | Random Forest | 1260 | 0.654 | 0.681 | 0.235 | 0.617 |

Warning: Native-row populations differ between methods; these results are secondary to Table 3.

# Table 5. Patient-Level Performance Summary

| Normalization | Domain | Classifier | N | Median | Q1 | Q3 | Min | Max |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Calibration median | All features | Extra Trees | 20 | 0.836 | 0.456 | 0.979 | 0.175 | 1.000 |
| Calibration median | All features | Logistic Regression | 20 | 0.414 | 0.234 | 0.698 | 0.000 | 1.000 |
| Calibration median | All features | Random Forest | 20 | 0.795 | 0.627 | 0.988 | 0.215 | 1.000 |
| Calibration median | Reduced features | Extra Trees | 20 | 0.801 | 0.557 | 0.942 | 0.321 | 1.000 |
| Calibration median | Reduced features | Logistic Regression | 20 | 0.381 | 0.144 | 0.649 | 0.000 | 0.989 |
| Calibration median | Reduced features | Random Forest | 20 | 0.851 | 0.611 | 0.993 | 0.217 | 1.000 |
| Calibration median-IQR | All features | Extra Trees | 20 | 0.922 | 0.685 | 1.000 | 0.277 | 1.000 |
| Calibration median-IQR | All features | Logistic Regression | 20 | 0.362 | 0.166 | 0.671 | 0.000 | 0.952 |
| Calibration median-IQR | All features | Random Forest | 20 | 0.845 | 0.615 | 1.000 | 0.337 | 1.000 |
| Calibration median-IQR | Reduced features | Extra Trees | 20 | 0.873 | 0.640 | 0.968 | 0.279 | 1.000 |
| Calibration median-IQR | Reduced features | Logistic Regression | 20 | 0.494 | 0.349 | 0.737 | 0.000 | 0.940 |
| Calibration median-IQR | Reduced features | Random Forest | 20 | 0.890 | 0.635 | 0.973 | 0.331 | 1.000 |
| Retrospective jackknife | All features | Extra Trees | 20 | 0.936 | 0.796 | 1.000 | 0.378 | 1.000 |
| Retrospective jackknife | All features | Logistic Regression | 20 | 0.416 | 0.107 | 0.641 | 0.000 | 1.000 |
| Retrospective jackknife | All features | Random Forest | 20 | 0.918 | 0.759 | 1.000 | 0.284 | 1.000 |
| Retrospective jackknife | Reduced features | Extra Trees | 20 | 0.953 | 0.799 | 0.985 | 0.346 | 1.000 |
| Retrospective jackknife | Reduced features | Logistic Regression | 20 | 0.427 | 0.198 | 0.663 | 0.000 | 1.000 |
| Retrospective jackknife | Reduced features | Random Forest | 20 | 0.966 | 0.857 | 0.987 | 0.473 | 1.000 |
| No normalization | All features | Extra Trees | 20 | 0.682 | 0.266 | 0.821 | 0.000 | 1.000 |
| No normalization | All features | Logistic Regression | 20 | 0.495 | 0.342 | 0.738 | 0.002 | 1.000 |
| No normalization | All features | Random Forest | 20 | 0.696 | 0.524 | 0.825 | 0.030 | 1.000 |
| No normalization | Reduced features | Extra Trees | 20 | 0.831 | 0.658 | 0.912 | 0.043 | 1.000 |
| No normalization | Reduced features | Logistic Regression | 20 | 0.471 | 0.207 | 0.725 | 0.000 | 0.958 |
| No normalization | Reduced features | Random Forest | 20 | 0.849 | 0.433 | 0.925 | 0.087 | 1.000 |

# Table 6. Directionality Relative to No Normalization

| Domain | Classifier | Comparison | Improved | Worsened | Unchanged | Inverted→correct | Always inverted |
| --- | --- | --- | --- | --- | --- | --- | --- |
| All features | Extra Trees | Retrospective jackknife | 14 | 5 | 1 | 7 | 0 |
| All features | Extra Trees | Calibration median | 12 | 7 | 1 | 5 | 0 |
| All features | Extra Trees | Calibration median-IQR | 13 | 6 | 1 | 6 | 0 |
| All features | Logistic Regression | Retrospective jackknife | 8 | 11 | 1 | 4 | 3 |
| All features | Logistic Regression | Calibration median | 6 | 13 | 1 | 2 | 3 |
| All features | Logistic Regression | Calibration median-IQR | 9 | 11 | 0 | 3 | 3 |
| All features | Random Forest | Retrospective jackknife | 15 | 4 | 1 | 3 | 0 |
| All features | Random Forest | Calibration median | 14 | 5 | 1 | 2 | 0 |
| All features | Random Forest | Calibration median-IQR | 13 | 5 | 2 | 2 | 0 |
| Reduced features | Extra Trees | Retrospective jackknife | 13 | 6 | 1 | 4 | 1 |
| Reduced features | Extra Trees | Calibration median | 9 | 11 | 0 | 3 | 1 |
| Reduced features | Extra Trees | Calibration median-IQR | 9 | 11 | 0 | 3 | 1 |
| Reduced features | Logistic Regression | Retrospective jackknife | 8 | 12 | 0 | 1 | 5 |
| Reduced features | Logistic Regression | Calibration median | 6 | 13 | 1 | 2 | 5 |
| Reduced features | Logistic Regression | Calibration median-IQR | 9 | 10 | 1 | 3 | 5 |
| Reduced features | Random Forest | Retrospective jackknife | 16 | 3 | 1 | 5 | 0 |
| Reduced features | Random Forest | Calibration median | 10 | 10 | 0 | 3 | 0 |
| Reduced features | Random Forest | Calibration median-IQR | 12 | 8 | 0 | 4 | 0 |

Figure 4. Per-patient change in ROC AUC relative to no normalization for all-feature Random Forest. Positive values indicate improvement.

Figure 5. Retrospective-to-calibration gap for the two tree ensembles. Fixed calibration recovers only part of the retrospective normalization benefit.

# 14. Patient-Level Heterogeneity

Patient-level analysis is essential because pooled metrics combine within-patient ranking with cross-patient score alignment. Jackknife tree-model gains were broadly distributed, particularly for reduced-feature Random Forest, where 16 of 20 patients improved and no patient remained inverted. Fixed calibration was less uniform. A pooled improvement could occur even when a majority of patients did not improve, as observed when reduced-feature Random Forest moved from median to median-IQR scaling.

Clinical-context examples show that large gains occurred in clean-sinus, atrial-fibrillation, and mechanically supported patients, while some paced patients already had strong unnormalized discrimination and worsened after centring. These are descriptive observations only; the frozen evidence does not include a prespecified subgroup or interaction analysis.

# 15. Calibration and IQR Fallback Findings

The fixed calibration set contained 200 labelled reference observations: 10 per patient. After exclusion, 1,060 observations remained for evaluation. Because row ordering was deterministic but not clinically time-verified, the method is best described as fixed limited labelled calibration rather than prospective calibration.

Median-IQR scaling required 3,240 fallback applications among 18,000 eligible patient-feature-fold operations, an overall proportion of 18%. Fallback affected all 20 patients and 20 of 35 features. fd_bandPower required fallback in every eligible application, Petrosian fractal dimension in 85%, variance in 60%, mean absolute value in 55%, and median in 50%.

Figure 6. Feature-level and patient-level median-IQR fallback proportions. Rates are derived from the frozen aggregate fallback counts and the documented design denominators.

# Table 7. Calibration and IQR Fallback Summary

| Item | Frozen result |
| --- | --- |
| Calibration patients | 20 |
| Rows per patient | 10 labelled reference rows |
| Total calibration rows | 200 |
| Evaluation rows after exclusion | 1,060 |
| Overall fallback applications | 3,240 |
| Eligible applications | 18,000 |
| Overall fallback proportion | 18.0% |
| Patients affected | 20 of 20 |
| Features affected | 20 of 35 |
| Highest feature fallback | fd_bandPower: 100% |

# Table 8. Clinical Cohort Context

| Characteristic | Count or summary |
| --- | --- |
| Patients | 20 |
| Median age | 69.5 years |
| Age IQR | 55.75-79 years |
| Men / women | 13 / 7 |
| Any atrial fibrillation | 6 |
| Pacing or ICD concern | 5 |
| Bundle branch block | 4 |
| Antiarrhythmic exposure | 5 |
| Mechanical circulatory support | 3 |
| Atypical/non-cardiogenic HI | 3 |
| Clean-sinus candidates | 4 |
| Outside strict HRV-valid subgroup | 8 |

Clinical characteristics are provided for cohort context. No frozen subgroup-performance analysis was available.

# 16. Main Scientific Interpretation

Patient-specific variation is a major part of the problem. High within-patient AUC but modest pooled AUC without normalization indicates that patient-specific feature offsets and scale differences impair cross-patient score alignment.

Patient adaptation can substantially improve nonlinear classifiers. Retrospective centring increased tree-model AUC by 0.185-0.280 and corrected directionality for many patients.

Retrospective normalization is optimistic. It uses extensive labelled information from the held-out patient and does not represent initial real-time calibration.

Limited calibration recovers only part of the benefit. The tree-model gap between jackknife and fixed calibration was 0.064-0.169.

Median-IQR scaling offers a trade-off. It improved pooled tree-model AUC but depended frequently on training-derived scale fallback and did not benefit every patient.

Aggregate performance is not patient-uniform. Pooled AUC and patient improvement counts must be reported together.

# 17. What the Study Does and Does Not Establish

## 17.1 Establishes

Patient-specific variation materially changes pooled tree-model performance.

Retrospective jackknife is an optimistic comparator of personalization potential.

Fixed labelled calibration recovers part of the benefit.

Patient responses are heterogeneous.

IQR fallback is frequent under a 10-row calibration sample.

## 17.2 Does not establish

Prospective real-time detection or future-event prediction.

Clinical readiness or bedside utility.

Performance using unlabelled calibration.

Verified chronological calibration.

External generalisability or formal statistical superiority.

Causal effects of rhythm, devices, medications, or diagnosis.

# 18. Limitations

Only 20 independent patients despite 1,260 repeated observations.

Clinically heterogeneous, enriched cohort with several potential HI mechanisms.

No external or prospective validation.

Calibration required labelled reference observations.

Calibration order was not verified as chronological.

Only 10 calibration rows were used; scale estimation was frequently unstable.

No frozen confidence intervals or paired statistical tests.

Upstream waveform processing, segmentation, labelling, and feature extraction were not present in the minimal repository.

The full feature set contains correlated or mathematically related quantities.

Fixed classifier specifications were not optimized.

Temporal dependence and possible window overlap could not be independently characterized.

The manifest recorded a non-clean working tree, and exact-environment full rerun was not completed during report assembly.

# 19. Implications for a Practical Monitoring System

A practical system would require a clearly defined, chronological calibration period that precedes prediction and represents a stable patient state. It would also need a policy for uncertain calibration labels, baseline drift, rhythm conversion, electrode changes, insufficient feature variation, and reversion to a non-personalized model when calibration quality is poor.

Clinical translation would require external validation, prospective data collection, explicit lead-time or detection definitions, threshold selection, alarm-burden analysis, subgroup and failure-mode evaluation, clinician usability assessment, and demonstration of impact on decisions or outcomes.

# 20. Proposed Manuscript Story

Haemodynamic instability is clinically important, while continuous invasive blood-pressure monitoring is not universally available. ECG is a routinely acquired continuous signal, but patient-specific physiology and devices create substantial variation in engineered feature distributions. The study asks whether patient adaptation can reduce this cross-patient misalignment and, critically, whether the apparent benefit survives when patient information is restricted to a small calibration set.

Without normalization, tree models often discriminated well within patients but showed modest pooled performance. Retrospective jackknife centring substantially improved discrimination, whereas fixed 10-row labelled calibration recovered only part of that benefit. Median-IQR scaling improved pooled calibration performance but required frequent fallback and had mixed patient-level effects. The manuscript contribution is therefore a patient-aware evaluation framework that distinguishes retrospective personalization potential from limited-calibration performance.

# 21. Candidate Manuscript Contributions

Leakage-safe LOSO comparison of four patient-normalization strategies.

Explicit separation of retrospective adaptation from fixed limited calibration.

Primary common-row evaluation holding row identity constant.

Patient-level directionality analysis of improved, worsened, and inverted patients.

Operational audit of median-IQR fallback under a small calibration sample.

# 22. Recommended Tables and Figures

The collaborator report includes eight tables and six figures. The main manuscript should retain the dataset/design table, primary common-row results, patient-level summary and directionality, plus Figures 1-5. Native-row details, the complete experimental matrix, detailed fallback diagnostics, and clinical-context overlays are better suited to supplementary material.

# 23. Conclusions

Patient adaptation can materially improve ECG-based classification of haemodynamic instability, but retrospective normalization should not be interpreted as prospective deployment performance.

The strongest evidence is the discrepancy between pooled and within-patient discrimination without normalization, the broad tree-model improvement under retrospective jackknife centring, and the incomplete recovery under fixed 10-row labelled calibration. Median-IQR scaling improved pooled tree-model discrimination but exposed an important practical limitation: reliable patient-specific dispersion could not always be estimated from a small calibration sample.

Future evaluation should prioritize chronologically acquired, minimally labelled or unlabelled calibration in independent and prospectively collected cohorts.

# 24. Technical Appendix

## 24.1 Authoritative artifacts

| Artifact | Role |
| --- | --- |
| data/FullFeatureDataset_Clean.csv | Canonical modelling dataset |
| results/frozen/manifest.json | Configuration, environment, provenance |
| results/frozen/oof_predictions.csv | Row-level OOF predictions |
| results/frozen/common_row_metrics.csv | Primary pooled results |
| results/frozen/native_row_metrics.csv | Secondary pooled results |
| results/frozen/common_row_patient_metrics.csv | Patient-level common-row metrics |
| results/frozen/patient_directionality.csv | Improved/worsened counts |
| results/frozen/calibration_rows.csv | Calibration identities |
| results/frozen/iqr_fallback_counts.csv | Fallback diagnostics |
| results/frozen/fold_summary.csv | Fold counts, selected features, hashes |

## 24.2 Frozen software and provenance

| Field | Frozen value |
| --- | --- |
| Python | 3.13.9 |
| NumPy | 2.3.5 |
| pandas | 2.3.3 |
| scikit-learn | 1.7.2 |
| Git branch | audit/student-code-review |
| Git commit | a58aa971a26986fb7188bdb613ab4c133f3dfae8 |
| Working tree clean | false |
| Run timestamp | 2026-07-30T15:56:50.169119+00:00 |

## 24.3 Selected checksums

| File | SHA-256 |
| --- | --- |
| FullFeatureDataset_Clean.csv | ef5a8c73034afa0ff92da1a5c31d133c1d847fe1894ba8c559bfd2cdab440796 |
| calibration_rows.csv | 17ac6131288d7502b4ea0862a96ab1130acdc54a243ae634955f1b66b6ae32fa |
| common_row_metrics.csv | 4b9cd0c9f1ef9f62cd6f7c5c4f0f654c2860dfed311d46a21aa753c52049c38d |
| common_row_patient_metrics.csv | f871b2e40d9784af7cdfda2efa6eea1ecdae3c13e3271fce8a22fdf5885a283e |
| fold_summary.csv | fb5fe60fca087d4cb8e9ecd8adcd4c12bd737899a89c2b7c9e10b981d5f1cc13 |
| iqr_fallback_counts.csv | 9b6fc6daa5e07ac0a687216330084eef5279fb874913dcdf0dab54573a981b98 |
| manifest.json | 88c554a0498f06396992a50c5706f2771f9104b0251512e9ac6fe58238bfeec2 |
| native_row_metrics.csv | 636d29e4503848288e43f2f450288774d1d98049bc58984f862d152825be6ade |
| native_row_patient_metrics.csv | ab37f2a699db2b2e9c02620b944785f82364596144d3fcfed7562a85d5738609 |
| oof_predictions.csv | 091e69a66b3f0c731b3a0d964d2dece1a04d63d10b6d5f74d664b3e7ffc29d70 |
| patient_directionality.csv | a9e1f698f869c849e610f5f66449ea124f21d8cbb708a6771d8b33d63eb2b4b1 |

## 24.4 Traceability matrix

| Claim | Primary source |
| --- | --- |
| 20 patients; 1,260 rows | Dataset, manifest |
| Common comparison used 1,060 identical rows | OOF predictions, common metrics |
| Best retrospective AUC 0.866 | Common metrics: jackknife/reduced/RF |
| Best fixed-calibration AUC 0.769 | Common metrics: median-IQR/all/ET |
| 16 patients improved in best retrospective configuration | Directionality summary |
| Retrospective-to-calibration gap 0.064-0.169 | Matched common-metric subtraction |
| Fallback proportion 18% | IQR fallback counts |
| Calibration chronology unverified | Calibration rows + identifier audit |
| Clinical heterogeneity present | Clinical summary and flags |

## 24.5 Audit qualification

The frozen evidence is sufficient to reconstruct the patient-normalization benchmark and its major numerical conclusions. It is not sufficient to independently audit the upstream waveform and event-labelling pipeline, establish prospective calibration, conduct formal clinical subgroup inference, or claim external clinical readiness.